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Supplemental Table 1:  Modifications to gradient and Multiple Reaction Monitoring (MRM) parameters within current LC-MS/MS method

A.

	Time (min)
	%A*
	%B**

	Initial
	28
	72

	5.00
	28
	72

	5.15
	2
	98

	6.15
	2
	98

	6.30
	28
	72

	7.0
	28
	72



	*Mobile Phase A:  Water + 0.1% formic acid, 2mM ammonium acetate
**Mobile Phase B:  Methanol + 0.1% formic acid, 2mM ammonium acetate




B.

	Analyte
	Parent Ion
	Daughter Ion
	Dwell time (s)
	Cone (V)
	Collision Energy (V)

	25(OH)D₃
	401.3
	159.0 (quant)
	0.200
	24
	28

	
	401.3
	365.4
	0.200
	22
	11

	d3-3epi-25(OH)D₃
	404.4
	162.0
	0.200
	21
	27

	d6-25(OH)D₃
	407.4
	159.1
	0.200
	20
	25

	25(OH)D₂
	413.3
	82.9
	0.200
	23
	22

	
	413.3
	355.3 (quant)
	0.200
	23
	11

	[13C]3-25(OH)D₂
	416.2
	355.2
	0.200
	17
	14




