Online Supporting Material

Supplemental Table 1:  Modifications to gradient and Multiple Reaction Monitoring (MRM) parameters within current LC-MS/MS method

A.

| Time (min) | %A* | %B** |
| --- | --- | --- |
| Initial | 28 | 72 |
| 5.00 | 28 | 72 |
| 5.15 | 2 | 98 |
| 6.15 | 2 | 98 |
| 6.30 | 28 | 72 |
| 7.0 | 28 | 72 |

*Mobile Phase A: Water + 0.1% formic acid, 2mM ammonium acetate **Mobile Phase B: Methanol + 0.1% formic acid, 2mM ammonium acetate

B.

| Analyte | Parent Ion | Daughter Ion | Dwell time (s) | Cone (V) | Collision Energy (V) |
| --- | --- | --- | --- | --- | --- |
| 25(OH)D₃ | 401.3 | 159.0 (quant) | 0.200 | 24 | 28 |
|  | 401.3 | 365.4 | 0.200 | 22 | 11 |
| d3-3epi-25(OH)D₃ | 404.4 | 162.0 | 0.200 | 21 | 27 |
| d6-25(OH)D₃ | 407.4 | 159.1 | 0.200 | 20 | 25 |
| 25(OH)D₂ | 413.3 | 82.9 | 0.200 | 23 | 22 |
|  | 413.3 | 355.3 (quant) | 0.200 | 23 | 11 |
| [13C]3-25(OH)D₂ | 416.2 | 355.2 | 0.200 | 17 | 14 |