## Reporting Template General
Participant Laboratory Information
Primary contact
Institution
Telephone number
E-mail
Assay description/characteristics
LC-MS/MS Methods
Instrumention description
Chromatography column (manufacturer, dimensions, particle size)
Chromatography conditions (briefly)
Internal Standards used
Transitions monitored for analytes and internal standards
Published method citation (if available)
Analyte information
|  |  | Reportable range (nmol/L) |  | Comments
Analytes reported from your method |  | Limit of detection (nmol/L) | Lower limit | Upper limit
24,25-dihydroxyvitamin D3
List any others
Internal Standards
d6-24 (R),25 dihydroxivitamin D3
Standards and Control Material Information
Analyte Standards | Source | Lot number | Purity Stated | Purity assessed | Comments
24,25-dihydroxyvitamin D3
Control Materials used (other than SRM 972a and SRM 2973 that were provided)
SRM 972a L1
DEQAS 616
DEQAS 618

## Reporting Template - Results
Sample Results
Date of Analysis
Measured values (nmol/L); 3 significant digits
Run order | Sample Code | Vial ID 1 | Vial ID 2 ( if necessary) | 24,25(OH)2D3 | Comments/Explanation of Any Error Messages
1 | SRM972aL2A-1 |  |  |  |  | :
2 | DS01-1
3 | DS02-1
4 | DS03-1
5 | DS04-1
6 | DS05-1
7 | DS06-1
8 | DS07-1
9 | DS08-1
10 | SRM2973A-1
11 | DS09-1
12 | DS10-1
13 | DS11-1
14 | DS12-1
15 | DS13-1
16 | DS14-1
17 | DS15-1
18 | DS16-1
19 | SRM972aL2B-1
20 | DS17-1
21 | DS18-1
22 | DS19-1
23 | DS20-1
24 | DS21-1
25 | DS22-1
26 | DS23-1
27 | DS24-1
28 | SRM2973B-1
29 | DS25-1
30 | DS26-1
31 | DS27-1
32 | DS28-1
33 | DS29-1
34 | DS30-1
35 | DS31-1
36 | DS32-1
37 | SRM972aL2C-1
38 | DS33-1
39 | DS34-1
40 | DS35-1
41 | DS36-1
42 | DS37-1
43 | DS38-1
44 | DS39-1
45 | DS40-1
46 | SRM2973C-1
47 | DS41-1
48 | DS42-1
49 | DS43-1
50 | DS44-1
51 | DS45-1
52 | DS46-1
53 | DS47-1
54 | DS48-1
55 | SRM972aL2D-1
56 | DS49-1
57 | DS50-1
58 | SRM2969A-1
59 | SRM2969B-1
60 | SRM2969C-1
61 | SRM2970A-1
62 | SRM2970B-1
63 | SRM2970C-1
64 | SRM2973D-1
65 | SRM972aL2A-2
66 | DS01-2
67 | DS02-2
68 | DS03-2
69 | DS04-2
70 | DS05-2
71 | DS06-2
72 | DS07-2
73 | DS08-2
74 | SRM2973A-2
75 | DS09-2
76 | DS10-2
77 | DS11-2
78 | DS12-2
79 | DS13-2
80 | DS14-2
81 | DS15-2
82 | DS16-2
83 | SRM972aL2B-2
84 | DS17-2
85 | DS18-2
86 | DS19-2
87 | DS20-2
88 | DS21-2
89 | DS22-2
90 | DS23-2
91 | DS24-2
92 | SRM2973B-2
93 | DS25-2
94 | DS26-2
95 | DS27-2
96 | DS28-2
97 | DS29-2
98 | DS30-2
99 | DS31-2
100 | DS32-2
101 | SRM972aL2C-2
102 | DS33-2
103 | DS34-2
104 | DS35-2
105 | DS36-2
106 | DS37-2
107 | DS38-2
108 | DS39-2
109 | DS40-2
110 | SRM2973C-2
111 | DS41-2
112 | DS42-2
113 | DS43-2
114 | DS44-2
115 | DS45-2
116 | DS46-2
117 | DS47-2
118 | DS48-2
119 | SRM972aL2D-2
120 | DS49-2
121 | DS50-2
122 | SRM2969A-2
123 | SRM2969B-2
124 | SRM2969C-2
125 | SRM2970A-2
126 | SRM2970B-2
127 | SRM2970C-2
128 | SRM2973D-2
129
130

## CDC data
CDC
Sample Code | Concentration 24,25(OH)2D3 (nmol/L)
| Replicate 1 | Replicate 2 | Mean | SD
DS01 | 0.877 | 0.923 | 0.9 | 0.032527
DS02 | 0.762 | 0.739 | 0.7505 | 0.016263
DS03 | 0.551 | 0.658 | 0.6045 | 0.07566
DS04 | 13.1 | 13.4 | 13.25 | 0.21213
DS05 | 6.95 | 6.98 | 6.965 | 0.021213
DS06 | 11.3 | 11.2 | 11.25 | 0.070711
DS07 | 6.1 | 6.21 | 6.155 | 0.077782
DS08 | 6.62 | 6.83 | 6.725 | 0.14849
DS09 | 0.441 | 0.469 | 0.455 | 0.019799
DS10 | 9.5 | 9.39 | 9.445 | 0.077782
DS11 | 7.37 | 7.34 | 7.355 | 0.021213
DS12 | 6.92 | 6.9 | 6.91 | 0.014142
DS13 | 6.98 | 6.96 | 6.97 | 0.014142
DS14 | 3.01 | 3 | 3.005 | 0.0070711
DS15 | 6.47 | 6.35 | 6.41 | 0.084853
DS16 | 3.91 | 4 | 3.955 | 0.06364
DS17 | 1.57 | 1.37 | 1.47 | 0.14142
DS18 | 1.44 | 1.41 | 1.425 | 0.021213
DS19 | 8.54 | 8.4 | 8.47 | 0.098995
DS20 | 4.21 | 4.07 | 4.14 | 0.098995
DS21 | 4.46 | 4.44 | 4.45 | 0.014142
DS22 | 0.687 | 0.719 | 0.703 | 0.022627
DS23 | 0.485 | 0.498 | 0.4915 | 0.0091924
DS24 | 1.38 | 1.26 | 1.32 | 0.084853
DS25 | 2.89 | 2.88 | 2.885 | 0.0070711
DS26 | 1.73 | 1.71 | 1.72 | 0.014142
DS27 | 7.84 | 7.36 | 7.6 | 0.33941
DS28 | 2.76 | 2.58 | 2.67 | 0.12728
DS29 | 3.15 | 3.09 | 3.12 | 0.042426
DS30 | 7.02 | 6.6 | 6.81 | 0.29698
DS31 | 1.07 | 0.988 | 1.029 | 0.057983
DS32 | 0.795 | 0.806 | 0.8005 | 0.0077782
DS33 | 3.27 | 3.32 | 3.295 | 0.035355
DS34 | 0.602 | 0.61 | 0.606 | 0.0056569
DS35 | 4 | 3.9 | 3.95 | 0.070711
DS36 | 6 | 6.09 | 6.045 | 0.06364
DS37 | 13.8 | 14.6 | 14.2 | 0.56569
DS38 | 12.2 | 11.9 | 12.05 | 0.21213
DS39 | 0.734 | 0.725 | 0.7295 | 0.006364
DS40 | 2.81 | 2.83 | 2.82 | 0.014142
DS41 | 5.38 | 5.51 | 5.445 | 0.091924
DS42 | 1.77 | 1.66 | 1.715 | 0.077782
DS43 | 2.73 | 2.5 | 2.615 | 0.16263
DS44 | 9.35 | 9.2 | 9.275 | 0.10607
DS45 | 3.44 | 3.43 | 3.435 | 0.0070711
DS46 | 5 | 4.89 | 4.945 | 0.077782
DS47 | 15.2 | 14.4 | 14.8 | 0.56569
DS48 | 1.95 | 1.88 | 1.915 | 0.049497
DS49 | 9.14 | 8.88 | 9.01 | 0.18385
DS50 | 6.51 | 6.08 | 6.295 | 0.30406
SRM2969A | 1.38 | 1.41 | 1.395 | 0.021213
SRM2969B | 1.45 | 1.38 | 1.415 | 0.049497
SRM2969C | 1.55 | 1.33 | 1.44 | 0.15556
MEAN |  |  | 1.4167 | 0.022546
SRM2970A | 1.78 | 1.68 | 1.73 | 0.070711
SRM2970B | 1.73 | 1.62 | 1.675 | 0.077782
SRM2970C | 1.74 | 1.67 | 1.705 | 0.049497
MEAN |  |  | 1.7033 | 0.027538
SRM2973A | 7.31 | 7.23 | 7.27 | 0.056569
SRM2973B | 7.01 | 6.76 | 6.885 | 0.17678
SRM2973C | 7.31 | 7.35 | 7.33 | 0.028284
SRM2973D | 7.51 | 6.92 | 7.215 | 0.41719
MEAN |  |  | 7.175 | 0.19896
SRM972aL2A | 3.25 | 3.19 | 3.22 | 0.042426
SRM972aL2B | 3.21 | 3.22 | 3.215 | 0.0070711
SRM972aL2C | 3.26 | 3.2 | 3.23 | 0.042426
SRM972aL2D | 3.32 | 3.11 | 3.215 | 0.14849
MEAN |  |  | 3.22 | 0.0070711

## ICHT data
ICHT
Sample Code | Concentration 24,25(OH)2D3 (nmol/L)
| Replicate 1 | Replicate 2 | Mean | SD
DS01 | 1.13 | 1.55 | 1.34 | 0.29698
DS02 | 0.968 | 0.914 | 0.941 | 0.038184
DS03 | 0.939 | 0.864 | 0.9015 | 0.053033
DS04 | 16 | 16.1 | 16.05 | 0.070711
DS05 | 8.42 | 8.43 | 8.425 | 0.0070711
DS06 | 12.9 | 13.2 | 13.05 | 0.21213
DS07 | 7.68 | 7.74 | 7.71 | 0.042426
DS08 | 8 | 7.53 | 7.765 | 0.33234
DS09 | 0.781 | 0.522 | 0.6515 | 0.18314
DS10 | 11.1 | 11.3 | 11.2 | 0.14142
DS11 | 8.61 | 8.54 | 8.575 | 0.049497
DS12 | 8.19 | 8.73 | 8.46 | 0.38184
DS13 | 8.06 | 8.12 | 8.09 | 0.042426
DS14 | 3.92 | 3.85 | 3.885 | 0.049497
DS15 | 7.43 | 7.58 | 7.505 | 0.10607
DS16 | 4.93 | 4.88 | 4.905 | 0.035355
DS17 | 1.95 | 2 | 1.975 | 0.035355
DS18 | 1.7 | 1.73 | 1.715 | 0.021213
DS19 | 10.4 | 10.3 | 10.35 | 0.070711
DS20 | 5.13 | 5.17 | 5.15 | 0.028284
DS21 | 5.31 | 5.58 | 5.445 | 0.19092
DS22 | 0.966 | 0.949 | 0.9575 | 0.012021
DS23 | 0.763 | 0.829 | 0.796 | 0.046669
DS24 | 1.75 | 1.62 | 1.685 | 0.091924
DS25 | 3.73 | 3.84 | 3.785 | 0.077782
DS26 | 1.83 | 1.7 | 1.765 | 0.091924
DS27 | 8.9 | 9.14 | 9.02 | 0.16971
DS28 | 3.43 | 3.53 | 3.48 | 0.070711
DS29 | 3.59 | 3.89 | 3.74 | 0.21213
DS30 | 8.29 | 8.21 | 8.25 | 0.056569
DS31 | 1.47 | 1.33 | 1.4 | 0.098995
DS32 | 0.98 | 0.949 | 0.9645 | 0.02192
DS33 | 4.27 | 4.2 | 4.235 | 0.049497
DS34 | 0.603 | 0.647 | 0.625 | 0.031113
DS35 | 4.91 | 4.9 | 4.905 | 0.0070711
DS36 | 7.14 | 7.04 | 7.09 | 0.070711
DS37 | 16.2 | 16.1 | 16.15 | 0.070711
DS38 | 14.3 | 14.5 | 14.4 | 0.14142
DS39 | 0.966 | 0.834 | 0.9 | 0.093338
DS40 | 3.51 | 3.51 | 3.51 | 0
DS41 | 6.48 | 6.4 | 6.44 | 0.056569
DS42 | 2.23 | 2.2 | 2.215 | 0.021213
DS43 | 3.16 | 3.3 | 3.23 | 0.098995
DS44 | 11.2 | 11.3 | 11.25 | 0.070711
DS45 | 4.31 | 4.29 | 4.3 | 0.014142
DS46 | 5.99 | 5.86 | 5.925 | 0.091924
DS47 | 17.7 | 18 | 17.85 | 0.21213
DS48 | 2.58 | 2.44 | 2.51 | 0.098995
DS49 | 11 | 11.1 | 11.05 | 0.070711
DS50 | 7.9 | 7.77 | 7.835 | 0.091924
SRM2973A | 8.31 | 8.66 | 8.485 | 0.24749
SRM2973B | 8.46 | 8.28 | 8.37 | 0.12728
SRM2973C | 8.26 | 8.25 | 8.255 | 0.0070711
SRM2973D | 8.48 | 8.37 | 8.425 | 0.077782
Mean |  |  | 8.3837 | 0.10122
SRM972aL2A | 4.27 | 3.94 | 4.105 | 0.23335
SRM972aL2B | 3.93 | 4.04 | 3.985 | 0.077782
SRM972aL2C | 3.91 | 3.75 | 3.83 | 0.11314
SRM972aL2D | 4.03 | 4.31 | 4.17 | 0.19799
Mean |  |  | 4.0225 | 0.072342

## U College Cork data
U College Cork
Sample Code | Concentration 24,25(OH)2D3 (nmol/L)
| Replicate 1 | Replicate 2 | Mean | SD
DS01 | <0.65 | <0.65
DS02 | <0.65 | <0.65
DS03 | <0.65 | <0.65
DS04 | 15.4175 | 17.3935 | 16.4055 | 1.3972
DS05 | 7.1357 | 8.3557 | 7.7457 | 0.86262
DS06 | 11.4141 | 15.4175 | 13.4158 | 2.8309
DS07 | 5.7955 | 6.1907 | 5.9931 | 0.27944
DS08 | 7.1186 | 9.0086 | 8.0636 | 1.3365
DS09 | <0.65 | <0.65
DS10 | 11.727 | 10.8127 | 11.2699 | 0.6465
DS11 | 7.9261 | 9.6443 | 8.7852 | 1.215
DS12 | 6.9639 | 8.4416 | 7.7027 | 1.0449
DS13 | 9.1976 | 8.0636 | 8.6306 | 0.80187
DS14 | 3.5619 | 3.2182 | 3.39 | 0.24299
DS15 | 7.3419 | 7.4107 | 7.3763 | 0.048598
DS16 | 5.3144 | 5.0567 | 5.1856 | 0.18224
DS17 | 1.6031 | 1.5687 | 1.5859 | 0.024299
DS18 | 1.4656 | 1.2251 | 1.3454 | 0.17009
DS19 | 8.4759 | 9.9708 | 9.2234 | 1.057
DS20 | 4.9708 | 4.2835 | 4.6271 | 0.48598
DS21 | 3.5619 | 5.4863 | 4.5241 | 1.3608
DS22 | <0.65 | <0.65
DS23 | <0.65 | <0.65
DS24 | 0.8299 | 0.69244 | 0.76117 | 0.097197
DS25 | 2.46 | 2.793 | 2.6265 | 0.23547
DS26 | 2.08 | 1.8093 | 1.9446 | 0.19143
DS27 |  | 8.8883 | 8.8883
DS28 | 2.5309 | 2.6856 | 2.6082 | 0.10935
DS29 | 4.1976 | 2.3763 | 3.2869 | 1.2879
DS30 | 6.689 | 6.0704 | 6.3797 | 0.43739
DS31 | 1.2423 | 1.2251 | 1.2337 | 0.01215
DS32 | <0.65 | <0.65
DS33 | 3.0464 | 6.1564 | 4.6014 | 2.1991
DS34 | <0.65 | <0.65
DS35 | 4.8677 | 3.4759 | 4.1718 | 0.98412
DS36 | 7.5825 | 8.6478 | 8.1151 | 0.75328
DS37 | 18.7165 | 15.0395 | 16.878 | 2.6
DS38 | 14.3694 | 15.413 | 14.8912 | 0.73793
DS39 | <0.65 | <0.65
DS40 | 2.4966 | 2.3591 | 2.4278 | 0.097197
DS41 | 6.3969 | 7.2388 | 6.8179 | 0.59533
DS42 | 1.6203 | 1.2766 | 1.4485 | 0.24299
DS43 | 2.1357 | 2.0326 | 2.0842 | 0.072898
DS44 | 10.3488 | 9.2491 | 9.799 | 0.77757
DS45 | 4.6443 | 3.6478 | 4.146 | 0.70468
DS46 | 5.7612 | 6.898 | 6.3296 | 0.80386
DS47 | 17.9261 | 17.7543 | 17.8402 | 0.1215
DS48 | 1.2766 | 1.7577 | 1.5172 | 0.34019
DS49 | 11.9296 | 8.7509 | 10.3402 | 2.2477
DS50 | 7.6856 | 8.6821 | 8.1838 | 0.70468
SRM2969A | 1.59 | 1.14 | 1.365 | 0.3182
SRM2969B | 0.85 | 0.92 | 0.885 | 0.049497
SRM2969C | 1.21 | 1.26 | 1.235 | 0.035355
Mean |  |  | 1.1617 | 0.24826
SRM2970A | 3.012 | 2.067 | 2.5395 | 0.66823
SRM2970B | 1.6031 | 1.0361 | 1.3196 | 0.40094
SRM2970C | 1.5687 | 2.0498 | 1.8093 | 0.34019
Mean |  |  | 1.8895 | 0.61391
SRM2973A | 7.1186 | 8.5275 | 7.823 | 0.99627
SRM2973B | 7.5137 | 7.4622 | 7.488 | 0.036449
SRM2973C | 7.9777 | 8.2698 | 8.1237 | 0.20654
SRM2973D | 9.5412 | 8.4244 | 8.9828 | 0.78972
Mean |  |  | 8.1982 | 0.7502
SRM972aL2A | 3.3729 | 3.768 | 3.5704 | 0.27944
SRM972aL2B | 3.9399 | 3.8711 | 3.9055 | 0.048598
SRM972aL2C | 3.0464 | 4.3866 | 3.7165 | 0.94767
SRM972aL2D | 4.4553 | 3.4588 | 3.957 | 0.70468
Mean |  |  | 3.8597 | 0.12665

## U of Liege Data
Concnetration 24,25(OH)2D3 (nmol/L)
Sample Code | Replicate 1 | Replicate 2 | Mean | SD
DS01 | 1.0995 | 0.95816 | 1.0288 | 0.099965
DS02 | 0.89287 | 0.97328 | 0.93308 | 0.056856
DS03 | 0.72558 | 0.80887 | 0.76722 | 0.058893
DS04 | 12.363 | 13.299 | 12.831 | 0.66191
DS05 | 6.5655 | 7.3187 | 6.9421 | 0.53258
DS06 | 11.3717 | 11.8829 | 11.6273 | 0.3615
DS07 | 6.4733 | 6.3048 | 6.3891 | 0.11914
DS08 | 6.9059 | 6.7578 | 6.8318 | 0.10472
DS09 | <0.624 | <0.624
DS10 | 9.391 | 9.5209 | 9.4559 | 0.091818
DS11 | 6.8079 | 6.8655 | 6.8367 | 0.040733
DS12 | 6.8929 | 7.2059 | 7.0494 | 0.22131
DS13 | 7.5174 | 7.2628 | 7.3901 | 0.18007
DS14 | 3.0936 | 3.192 | 3.1428 | 0.069585
DS15 | 6.1239 | 5.6822 | 5.9031 | 0.31228
DS16 | 3.8605 | 4.1005 | 3.9805 | 0.16972
DS17 | 1.6674 | 1.5071 | 1.5873 | 0.11337
DS18 | 1.4795 | 1.4912 | 1.4854 | 0.0083163
DS19 | 8.7802 | 8.5164 | 8.6483 | 0.18652
DS20 | 3.9687 | 4.2752 | 4.122 | 0.21673
DS21 | 4.3223 | 4.4987 | 4.4105 | 0.12474
DS22 | 0.79375 | 0.75918 | 0.77646 | 0.02444
DS23 | <0.624 | <0.624
DS24 | 1.4461 | 1.5729 | 1.5095 | 0.089612
DS25 | 2.9611 | 3.025 | 2.993 | 0.045145
DS26 | 1.8635 | 1.9355 | 1.8995 | 0.050916
DS27 | 7.5088 | 7.4128 | 7.4608 | 0.067888
DS28 | 2.3186 | 2.609 | 2.4638 | 0.20536
DS29 | 3.259 | 3.2045 | 3.2317 | 0.038526
DS30 | 6.9577 | 6.7933 | 6.8755 | 0.11626
DS31 | 0.98696 | 0.95552 | 0.97124 | 0.022233
DS32 | 0.97424 | 0.93008 | 0.95216 | 0.031228
DS33 | 3.728 | 3.5401 | 3.634 | 0.13289
DS34 | 0.75054 | 0.54653 | 0.64853 | 0.14426
DS35 | 4.1919 | 4.2846 | 4.2383 | 0.065512
DS36 | 6.2782 | 6.0924 | 6.1853 | 0.13136
DS37 | 14.8162 | 14.3568 | 14.5865 | 0.32484
DS38 | 11.6935 | 12.9872 | 12.3404 | 0.91479
DS39 | 0.79927 | 0.67758 | 0.73842 | 0.086048
DS40 | 2.8087 | 2.9333 | 2.871 | 0.088085
DS41 | 5.4441 | 5.6289 | 5.5365 | 0.13068
DS42 | 1.8179 | 1.8434 | 1.8306 | 0.01799
DS43 | 2.6157 | 2.5915 | 2.6036 | 0.017142
DS44 | 9.6824 | 9.6478 | 9.6651 | 0.02444
DS45 | 3.8312 | 3.9318 | 3.8815 | 0.071113
DS46 | 4.57 | 4.8647 | 4.7174 | 0.20842
DS47 | 15.6639 | 16.474 | 16.069 | 0.5728
DS48 | 1.895 | 1.9288 | 1.9119 | 0.02393
DS49 | 8.9215 | 9.4213 | 9.1714 | 0.35336
DS50 | 6.9047 | 6.7863 | 6.8455 | 0.083672
SRM2969A | 1.2968 | 1.3499 | 1.3234 | 0.037508
SRM2969B | 1.4466 | 1.5126 | 1.4796 | 0.046673
SRM2969C | 1.5186 | 1.5104 | 1.5145 | 0.0057705
MEAN |  |  | 1.4392 | 0.1018
SRM2970A | 1.6194 | 1.7677 | 1.6936 | 0.10489
SRM2970B | 1.7111 | 1.9761 | 1.8436 | 0.18737
SRM2970C | 1.7641 | 1.863 | 1.8136 | 0.069925
MEAN |  |  | 1.7836 | 0.079379
SRM2973A | 7.3453 | 7.8187 | 7.582 | 0.33469
SRM2973B | 7.2397 | 7.6893 | 7.4645 | 0.31789
SRM2973C | 7.5779 | 7.4363 | 7.5071 | 0.10013
SRM2973D | 7.8575 | 7.5006 | 7.6791 | 0.25237
MEAN |  |  | 7.5582 | 0.094106
SRM972aL2A | 3.1486 | 3.5189 | 3.3338 | 0.26188
SRM972aL2B | 3.3879 | 3.4121 | 3.4 | 0.017142
SRM972aL2C | 3.5473 | 3.4796 | 3.5134 | 0.047861
SRM972aL2D | 3.42 | 3.2028 | 3.3114 | 0.1536
MEAN |  |  | 3.3897 | 0.090675

## U of WA data
Univ of WA
Sample Code | Concentration 24,25(OH)2D3 (nmol/L)
| Replicate 1 | Replicate 2 | Mean | SD
DS01 | 1.0771 | 0.99475 | 1.0359 | 0.058243
DS02 | 0.83213 | 0.85958 | 0.84586 | 0.019414
DS03 | 0.71174 | 0.81101 | 0.76138 | 0.07019
DS04 | 13.578 | 12.522 | 13.05 | 0.7467
DS05 | 6.845 | 7.4723 | 7.1586 | 0.44354
DS06 | 11.3245 | 11.8061 | 11.5653 | 0.3405
DS07 | 6.5197 | 7.1386 | 6.8292 | 0.43757
DS08 | 7.1787 | 9.6518 | 8.4153 | 1.7488
DS09 | 0.64838 | 0.63149 | 0.63994 | 0.011947
DS10 | 9.6666 | 9.523 | 9.5948 | 0.10155
DS11 | 7.1512 | 7.7743 | 7.4628 | 0.44056
DS12 | 6.7964 | 6.4437 | 6.6201 | 0.2494
DS13 | 6.5155 | 5.8967 | 6.2061 | 0.43757
DS14 | 3.3919 | 2.7097 | 3.0508 | 0.48237
DS15 | 6.1523 | 6.0931 | 6.1227 | 0.041815
DS16 | 4.11 | 3.9262 | 4.0181 | 0.12993
DS17 | 1.5185 | 1.4995 | 1.509 | 0.013441
DS18 | 1.3306 | 1.4277 | 1.3791 | 0.068697
DS19 | 8.7521 | 8.7606 | 8.7564 | 0.0059736
DS20 | 4.5366 | 4.0339 | 4.2852 | 0.35543
DS21 | 4.5302 | 4.4204 | 4.4753 | 0.077657
DS22 | 0.74131 | 1.0074 | 0.87437 | 0.18817
DS23 | 0.5977 | 0.69696 | 0.64733 | 0.07019
DS24 | 1.6854 | 1.358 | 1.5217 | 0.23148
DS25 | 2.7583 | 3.1215 | 2.9399 | 0.25687
DS26 | 1.8691 | 2.0191 | 1.9441 | 0.10603
DS27 | 6.9696 | 7.2948 | 7.1322 | 0.22999
DS28 | 2.6337 | 2.8681 | 2.7509 | 0.16577
DS29 | 3.2166 | 2.9124 | 3.0645 | 0.21505
DS30 | 6.8387 | 5.9347 | 6.3867 | 0.63918
DS31 | 1.0729 | 1.0412 | 1.0571 | 0.022401
DS32 | 0.98208 | 0.99686 | 0.98947 | 0.010454
DS33 | 3.6854 | 3.4806 | 3.583 | 0.14486
DS34 | 0.50477 | 0.71174 | 0.60826 | 0.14635
DS35 | 3.9896 | 4.0255 | 4.0075 | 0.025388
DS36 | 5.9094 | 6.7288 | 6.3191 | 0.57944
DS37 | 13.7702 | 13.4408 | 13.6055 | 0.23297
DS38 | 11.7807 | 11.7343 | 11.7575 | 0.032855
DS39 | 0.82579 | 0.70118 | 0.76349 | 0.088111
DS40 | 3.3306 | 2.7435 | 3.0371 | 0.41517
DS41 | 7.3836 | 4.8576 | 6.1206 | 1.7861
DS42 | 1.7551 | 2.0275 | 1.8913 | 0.19265
DS43 | 3.1068 | 2.999 | 3.0529 | 0.076164
DS44 | 9.7363 | 8.9485 | 9.3424 | 0.55704
DS45 | 3.846 | 3.7044 | 3.7752 | 0.10006
DS46 | 5.1934 | 5.0815 | 5.1374 | 0.079151
DS47 | 15.1958 | 14.8241 | 15.01 | 0.26284
DS48 | 1.7868 | 1.9705 | 1.8786 | 0.12993
DS49 | 9.8039 | 9.2675 | 9.5357 | 0.37933
DS50 | 6.7901 | 7.0435 | 6.9168 | 0.17921
SRM 2969A | 1.5544 | 1.6664 | 1.6104 | 0.079151
SRM 2969B | 1.7213 | 1.622 | 1.6716 | 0.07019
SRM 2969C | 1.7171 | 1.7635 | 1.7403 | 0.032855
Mean | 1.6643 | 1.684 | 1.6741 | 0.013938
SRM2970A | 1.9874 | 1.7234 | 1.8554 | 0.18668
SRM2970B | 1.9473 | 1.6262 | 1.7868 | 0.227
SRM2970C | 2.0402 | 1.7952 | 1.9177 | 0.17324
Mean | 1.9916 | 1.7149 | 1.8533 | 0.19564
SRM 2973A | 7.542 | 7.5736 | 7.5578 | 0.022401
SRM 2973B | 7.3054 | 7.1724 | 7.2389 | 0.094085
SRM 2973C |  | 7.1914 | 7.1914
SRM 2973D | 7.7658 | 7.3603 | 7.5631 | 0.28673
Mean | 7.5356 | 7.2413 | 7.3311 | 0.20808
SRM 972aL2A | 3.2102 | 3.6685 | 3.4394 | 0.32407
SRM 972aL2B | 3.432 | 3.5355 | 3.4837 | 0.073177
SRM 972aL2C | 3.1194 | 3.2356 | 3.1775 | 0.082138
SRM 972aL2D | 3.7446 | 3.6052 | 3.6749 | 0.098565
Mean | 3.432 | 3.4588 | 3.4454 | 0.018917

## Summary ALL DATA
Interlaboratory Comparison for Determination of 24,25(OH)2D3 - RESULTS
| NIST |  | CDC | [5 empty] | University of Washington | [5 empty] | University of Liege | [5 empty] | Univeristy College Cork RECAL | [5 empty] | Imperial College Healthcare NHS Trust (RECAL)
Donor Serum | Mean (nmol/L) (n=4) | SD | Mean of replicates | SD | Bias vs NIST | % Bias | ABS bias vs NIST | % ABS Bias | Mean of replicates | SD | Bias vs NIST | % Bias | ABS bias vs NIST | % ABS Bias | Mean of replicates | SD | Bias vs NIST | % Bias | ABS bias vs NIST | % ABS Bias | Mean of replicates | SD | Bias vs NIST | % Bias | ABS % bias vs NIST | % ABS Bias | Mean of replicates | SD | Bias vs NIST |  | % Bias | ABS % bias vs NIST | % ABS Bias
Log
Bias
1 | 0.8928 | 0.027253 | 0.9 | 0.032527 | 0.0072023 | 0.80671 | 0.0072023 | 0.80671 | 1.0359 | 0.058243 | 0.14314 | 16.0326 | 0.14314 | 16.0326 | 1.0288 | 0.099965 | 0.13605 | 15.2384 | 0.13605 | 15.2384 | [6 empty] | 1.34 | 0.29698 | 0.4472 | -0.3495 | 50.09 | 0.4472 | 50.09
2 | 0.75847 | 0.015611 | 0.7505 | 0.016263 | -0.0079663 | -1.0503 | 0.0079663 | 1.0503 | 0.84586 | 0.019414 | 0.08739 | 11.5219 | 0.08739 | 11.5219 | 0.93308 | 0.056856 | 0.17461 | 23.0216 | 0.17461 | 23.0216 | [6 empty] | 0.941 | 0.038184 | 0.18253 | -0.73866 | 24.0662 | 0.18253 | 24.0662
3 | 0.64693 | 0.029474 | 0.6045 | 0.07566 | -0.042433 | -6.5591 | 0.042433 | 6.5591 | 0.76138 | 0.07019 | 0.11444 | 17.6901 | 0.11444 | 17.6901 | 0.76722 | 0.058893 | 0.12029 | 18.594 | 0.12029 | 18.594 | [6 empty] | 0.9015 | 0.053033 | 0.25457 | -0.5942 | 39.3498 | 0.25457 | 39.3498
4 | 14.3837 | 0.14404 | 13.25 | 0.21213 | -1.1337 | -7.8817 | 1.1337 | 7.8817 | 13.05 | 0.7467 | -1.3336 | -9.2719 | 1.3336 | 9.2719 | 12.831 | 0.66191 | -1.5527 | -10.7948 | 1.5527 | 10.7948 | 16.4055 | 1.3972 | 2.0218 | 14.0563 | 2.0218 | 14.0563 | 16.05 | 0.070711 | 1.6663 | 0.22176 | 11.5848 | 1.6663 | 11.5848
5 | 7.0303 | 0.065966 | 6.965 | 0.021213 | -0.065274 | -0.92847 | 0.065274 | 0.92847 | 7.1586 | 0.44354 | 0.12835 | 1.8257 | 0.12835 | 1.8257 | 6.9421 | 0.53258 | -0.088176 | -1.2542 | 0.088176 | 1.2542 | 7.7457 | 0.86262 | 0.71543 | 10.1764 | 0.71543 | 10.1764 | 8.425 | 0.0070711 | 1.3947 | 0.14449 | 19.8389 | 1.3947 | 19.8389
6 | 10.9212 | 0.072962 | 11.25 | 0.070711 | 0.32883 | 3.011 | 0.32883 | 3.011 | 11.5653 | 0.3405 | 0.64414 | 5.8981 | 0.64414 | 5.8981 | 11.6273 | 0.3615 | 0.70612 | 6.4656 | 0.70612 | 6.4656 | 13.4158 | 2.8309 | 2.4946 | 22.8422 | 2.4946 | 22.8422 | 13.05 | 0.21213 | 2.1288 | 0.32814 | 19.4927 | 2.1288 | 19.4927
7 | 6.3194 | 0.070479 | 6.155 | 0.077782 | -0.16441 | -2.6017 | 0.16441 | 2.6017 | 6.8292 | 0.43757 | 0.50974 | 8.0663 | 0.50974 | 8.0663 | 6.3891 | 0.11914 | 0.069682 | 1.1027 | 0.069682 | 1.1027 | 5.9931 | 0.27944 | -0.32628 | -5.1632 | 0.32628 | 5.1632 | 7.71 | 0.042426 | 1.3906 | 0.1432 | 22.0051 | 1.3906 | 22.0051
8 | 6.4704 | 0.078248 | 6.725 | 0.14849 | 0.25461 | 3.935 | 0.25461 | 3.935 | 8.4153 | 1.7488 | 1.9449 | 30.0581 | 1.9449 | 30.0581 | 6.8318 | 0.10472 | 0.36142 | 5.5858 | 0.36142 | 5.5858 | 8.0636 | 1.3365 | 1.5932 | 24.6227 | 1.5932 | 24.6227 | 7.765 | 0.33234 | 1.2946 | 0.11214 | 20.0083 | 1.2946 | 20.0083
9 | 0.47899 | 0.010026 | 0.455 | 0.019799 | -0.023989 | -5.0082 | 0.023989 | 5.0082 | 0.63994 | 0.011947 | 0.16095 | 33.6015 | 0.16095 | 33.6015 | [12 empty] | 0.6515 | 0.18314 | 0.17251 | -0.76318 | 36.0157 | 0.17251 | 36.0157
10 | 9.2579 | 0.072413 | 9.445 | 0.077782 | 0.18706 | 2.0205 | 0.18706 | 2.0205 | 9.5948 | 0.10155 | 0.33688 | 3.6388 | 0.33688 | 3.6388 | 9.4559 | 0.091818 | 0.19801 | 2.1388 | 0.19801 | 2.1388 | 11.2699 | 0.6465 | 2.0119 | 21.7318 | 2.0119 | 21.7318 | 11.2 | 0.14142 | 1.9421 | 0.28826 | 20.9772 | 1.9421 | 20.9772
11 | 7.2394 | 0.1475 | 7.355 | 0.021213 | 0.11559 | 1.5967 | 0.11559 | 1.5967 | 7.4628 | 0.44056 | 0.22335 | 3.0851 | 0.22335 | 3.0851 | 6.8367 | 0.040733 | -0.40268 | -5.5623 | 0.40268 | 5.5623 | 8.7852 | 1.215 | 1.5458 | 21.3528 | 1.5458 | 21.3528 | 8.575 | 0.049497 | 1.3356 | 0.12567 | 18.4489
12 | 7.3049 | 0.16476 | 6.91 | 0.014142 | -0.39489 | -5.4058 | 0.39489 | 5.4058 | 6.6201 | 0.2494 | -0.68482 | -9.3748 | 0.68482 | 9.3748 | 7.0494 | 0.22131 | -0.2555 | -3.4976 | 0.2555 | 3.4976 | 7.7027 | 1.0449 | 0.39786 | 5.4465 | 0.39786 | 5.4465 | 8.46 | 0.38184 | 1.1551 | 0.062625 | 15.8129 | 1.1551 | 15.8129
13 | 6.9466 | 0.12163 | 6.97 | 0.014142 | 0.023354 | 0.33619 | 0.023354 | 0.33619 | 6.2061 | 0.43757 | -0.74053 | -10.6603 | 0.74053 | 10.6603 | 7.3901 | 0.18007 | 0.44345 | 6.3837 | 0.44345 | 6.3837 | 8.6306 | 0.80187 | 1.6839 | 24.241 | 1.6839 | 24.241 | 8.09 | 0.042426 | 1.1434 | 0.058181 | 16.4591 | 1.1434 | 16.4591
14 | 2.9382 | 0.0062378 | 3.005 | 0.0070711 | 0.066798 | 2.2734 | 0.066798 | 2.2734 | 3.0508 | 0.48237 | 0.11258 | 3.8317 | 0.11258 | 3.8317 | 3.1428 | 0.069585 | 0.20462 | 6.9641 | 0.20462 | 6.9641 | 3.39 | 0.24299 | 0.45183 | 15.3779 | 0.45183 | 15.3779 | 3.885 | 0.049497 | 0.9468 | -0.023743 | 32.2237 | 0.9468 | 32.2237
15 | 6.0615 | 0.071752 | 6.41 | 0.084853 | 0.34848 | 5.7491 | 0.34848 | 5.7491 | 6.1227 | 0.041815 | 0.061168 | 1.0091 | 0.061168 | 1.0091 | 5.9031 | 0.31228 | -0.15847 | -2.6143 | 0.15847 | 2.6143 | 7.3763 | 0.048598 | 1.3148 | 21.6904 | 1.3148 | 21.6904 | 7.505 | 0.10607 | 1.4435 | 0.15941 | 23.8138
16 | 4.5356 | 0.08507 | 3.955 | 0.06364 | -0.58061 | -12.8011 | 0.58061 | 12.8011 | 4.0181 | 0.12993 | -0.51753 | -11.4104 | 0.51753 | 11.4104 | 3.9805 | 0.16972 | -0.55512 | -12.2391 | 0.55512 | 12.2391 | 5.1856 | 0.18224 | 0.64996 | 14.3301 | 0.64996 | 14.3301 | 4.905 | 0.035355 | 0.36939 | -0.43252 | 8.1442 | 0.36939 | 8.1442
17 | 1.5881 | 0.029742 | 1.47 | 0.14142 | -0.11812 | -7.4377 | 0.11812 | 7.4377 | 1.509 | 0.013441 | -0.079096 | -4.9805 | 0.079096 | 4.9805 | 1.5873 | 0.11337 | -0.00086726 | -0.054609 | 0.00086726 | 0.054609 | 1.5859 | 0.024299 | -0.0022089 | -0.13909 | 0.0022089 | 0.13909 | 1.975 | 0.035355 | 0.38688 | -0.41242 | 24.3609 | 0.38688 | 24.3609
18 | 1.3756 | 0.01466 | 1.425 | 0.021213 | 0.049396 | 3.5909 | 0.049396 | 3.5909 | 1.3791 | 0.068697 | 0.0035321 | 0.25677 | 0.0035321 | 0.25677 | 1.4854 | 0.0083163 | 0.10976 | 7.979 | 0.10976 | 7.979 | 1.3454 | 0.17009 | -0.030243 | -2.1985 | 0.030243 | 2.1985 | 1.715 | 0.021213 | 0.3394 | -0.46929 | 24.6725 | 0.3394 | 24.6725
19 | 8.693 | 0.073714 | 8.47 | 0.098995 | -0.22301 | -2.5653 | 0.22301 | 2.5653 | 8.7564 | 0.0059736 | 0.063346 | 0.7287 | 0.063346 | 0.7287 | 8.6483 | 0.18652 | -0.044725 | -0.5145 | 0.044725 | 0.5145 | 9.2234 | 1.057 | 0.53036 | 6.101 | 0.53036 | 6.101 | 10.35 | 0.070711 | 1.657 | 0.21932 | 19.0612 | 1.657 | 19.0612
20 | 4.1748 | 0.026492 | 4.14 | 0.098995 | -0.034786 | -0.83323 | 0.034786 | 0.83323 | 4.2852 | 0.35543 | 0.11046 | 2.6459 | 0.11046 | 2.6459 | 4.122 | 0.21673 | -0.052802 | -1.2648 | 0.052802 | 1.2648 | 4.6271 | 0.48598 | 0.45236 | 10.8356 | 0.45236 | 10.8356 | 5.15 | 0.028284 | 0.97521 | -0.0109 | 23.3596 | 0.97521 | 23.3596
21 | 4.503 | 0.043877 | 4.45 | 0.014142 | -0.052957 | -1.176 | 0.052957 | 1.176 | 4.4753 | 0.077657 | -0.027629 | -0.61356 | 0.027629 | 0.61356 | 4.4105 | 0.12474 | -0.092469 | -2.0535 | 0.092469 | 2.0535 | 4.5241 | 1.3608 | 0.021098 | 0.46855 | 0.021098 | 0.46855 | 5.445 | 0.19092 | 0.94204 | -0.025929 | 20.9206 | 0.94204 | 20.9206
22 | 0.77606 | 0.028137 | 0.703 | 0.022627 | -0.073059 | -9.4142 | 0.073059 | 9.4142 | 0.87437 | 0.18817 | 0.098309 | 12.6677 | 0.098309 | 12.6677 | 0.77646 | 0.02444 | 0.00040528 | 0.052223 | 0.00040528 | 0.052223 | [6 empty] | 0.9575 | 0.012021 | 0.18144 | -0.74127 | 23.3797 | 0.18144 | 23.3797
23 | 0.52768 | 0.027768 | 0.4915 | 0.0091924 | -0.036177 | -6.8558 | 0.036177 | 6.8558 | 0.64733 | 0.07019 | 0.11965 | 22.6751 | 0.11965 | 22.6751 | [12 empty] | 0.796 | 0.046669 | 0.26832 | -0.57134 | 50.85 | 0.26832 | 50.85
24 | 1.5306 | 0.030105 | 1.32 | 0.084853 | -0.21061 | -13.7596 | 0.21061 | 13.7596 | 1.5217 | 0.23148 | -0.0089098 | -0.58211 | 0.0089098 | 0.58211 | 1.5095 | 0.089612 | -0.02112 | -1.3798 | 0.02112 | 1.3798 | [6 empty] | 1.685 | 0.091924 | 0.15439 | -0.81137 | 10.0871 | 0.15439 | 10.0871
25 | 2.9337 | 0.045847 | 2.885 | 0.0070711 | -0.048675 | -1.6592 | 0.048675 | 1.6592 | 2.9399 | 0.25687 | 0.0062291 | 0.21233 | 0.0062291 | 0.21233 | 2.993 | 0.045145 | 0.059374 | 2.0239 | 0.059374 | 2.0239 | 2.6265 | 0.23547 | -0.30717 | -10.4707 | 0.30717 | 10.4707 | 3.785 | 0.077782 | 0.85133 | -0.069905 | 29.0191
26 | 1.7865 | 0.014373 | 1.72 | 0.014142 | -0.066541 | -3.7246 | 0.066541 | 3.7246 | 1.9441 | 0.10603 | 0.15755 | 8.819 | 0.15755 | 8.819 | 1.8995 | 0.050916 | 0.11298 | 6.3238 | 0.11298 | 6.3238 | 1.9446 | 0.19143 | 0.1581 | 8.8494 | 0.1581 | 8.8494 | 1.765 | 0.091924 | -0.021541 |  | -1.2058 | 0.021541 | 1.2058
27 | 7.5152 | 0.058992 | 7.6 | 0.33941 | 0.084849 | 1.129 | 0.084849 | 1.129 | 7.1322 | 0.22999 | -0.38293 | -5.0954 | 0.38293 | 5.0954 | 7.4608 | 0.067888 | -0.054369 | -0.72346 | 0.054369 | 0.72346 | 8.8883 |  | 1.3732 | 18.272 | 1.3732 | 18.272 | 9.02 | 0.16971 | 1.5048 | 0.17749 | 20.0242 | 1.5048 | 20.0242
28 | 2.9978 | 0.034919 | 2.67 | 0.12728 | -0.32775 | -10.9333 | 0.32775 | 10.9333 | 2.7509 | 0.16577 | -0.24687 | -8.2353 | 0.24687 | 8.2353 | 2.4638 | 0.20536 | -0.53395 | -17.8116 | 0.53395 | 17.8116 | 2.6082 | 0.10935 | -0.38951 | -12.9932 | 0.38951 | 12.9932 | 3.48 | 0.070711 | 0.48225 | -0.31673 | 16.087 | 0.48225 | 16.087
29 | 3.1123 | 0.025425 | 3.12 | 0.042426 | 0.0077289 | 0.24834 | 0.0077289 | 0.24834 | 3.0645 | 0.21505 | -0.047759 | -1.5345 | 0.047759 | 1.5345 | 3.2317 | 0.038526 | 0.11948 | 3.8389 | 0.11948 | 3.8389 | 3.2869 | 1.2879 | 0.17467 | 5.6123 | 0.17467 | 5.6123 | 3.74 | 0.21213 | 0.62773 | -0.20223 | 20.1695 | 0.62773 | 20.1695
30 | 6.9015 | 0.080661 | 6.81 | 0.29698 | -0.091483 | -1.3255 | 0.091483 | 1.3255 | 6.3867 | 0.63918 | -0.51479 | -7.4592 | 0.51479 | 7.4592 | 6.8755 | 0.11626 | -0.02599 | -0.37658 | 0.02599 | 0.37658 | 6.3797 | 0.43739 | -0.52176 | -7.5601 | 0.52176 | 7.5601 | 8.25 | 0.056569 | 1.3485 | 0.12986 | 19.5395 | 1.3485 | 19.5395
31 | 1.2903 | 0.028929 | 1.029 | 0.057983 | -0.26129 | -20.2506 | 0.26129 | 20.2506 | 1.0571 | 0.022401 | -0.23324 | -18.0762 | 0.23324 | 18.0762 | 0.97124 | 0.022233 | -0.31905 | -24.727 | 0.31905 | 24.727 | 1.2337 | 0.01215 | -0.056615 | -4.3877 | 0.056615 | 4.3877 | 1.4 | 0.098995 | 0.10971 | -0.95976 | 8.5026 | 0.10971 | 8.5026
32 | 0.77797 | 0.021381 | 0.8005 | 0.0077782 | 0.022533 | 2.8964 | 0.022533 | 2.8964 | 0.98947 | 0.010454 | 0.2115 | 27.1869 | 0.2115 | 27.1869 | 0.95216 | 0.031228 | 0.17419 | 22.3907 | 0.17419 | 22.3907 | [6 empty] | 0.9645 | 0.02192 | 0.18653 | -0.72924 | 23.977 | 0.18653 | 23.977
33 | 3.4308 | 0.047615 | 3.295 | 0.035355 | -0.13582 | -3.9588 | 0.13582 | 3.9588 | 3.583 | 0.14486 | 0.15219 | 4.436 | 0.15219 | 4.436 | 3.634 | 0.13289 | 0.20321 | 5.9229 | 0.20321 | 5.9229 | 4.6014 | 2.1991 | 1.1706 | 34.1189 | 1.1706 | 34.1189 | 4.235 | 0.049497 | 0.80418 | -0.094646 | 23.44 | 0.80418 | 23.44
34 | 0.5757 | 0.034761 | 0.606 | 0.0056569 | 0.030301 | 5.2633 | 0.030301 | 5.2633 | 0.60826 | 0.14635 | 0.032557 | 5.6552 | 0.032557 | 5.6552 | 0.64853 | 0.14426 | 0.072835 | 12.6515 | 0.072835 | 12.6515 | [6 empty] | 0.625 | 0.031113 | 0.049301 | -1.3071 | 8.5636 | 0.049301 | 8.5636
35 | 3.6961 | 0.048411 | 3.95 | 0.070711 | 0.25391 | 6.8698 | 0.25391 | 6.8698 | 4.0075 | 0.025388 | 0.31143 | 8.426 | 0.31143 | 8.426 | 4.2383 | 0.065512 | 0.54219 | 14.6692 | 0.54219 | 14.6692 | 4.1718 | 0.98412 | 0.47573 | 12.8713 | 0.47573 | 12.8713 | 4.905 | 0.0070711 | 1.2089 | 0.082395 | 32.7079 | 1.2089 | 32.7079
36 | 6.034 | 0.13864 | 6.045 | 0.06364 | 0.011003 | 0.18235 | 0.011003 | 0.18235 | 6.3191 | 0.57944 | 0.28511 | 4.725 | 0.28511 | 4.725 | 6.1853 | 0.13136 | 0.15132 | 2.5078 | 0.15132 | 2.5078 | 8.1151 | 0.75328 | 2.0811 | 34.49 | 2.0811 | 34.49 | 7.09 | 0.070711 | 1.056 | 0.023665 | 17.5009 | 1.056 | 17.5009
37 | 13.8141 | 0.075907 | 14.2 | 0.56569 | 0.38591 | 2.7936 | 0.38591 | 2.7936 | 13.6055 | 0.23297 | -0.20858 | -1.5099 | 0.20858 | 1.5099 | 14.5865 | 0.32484 | 0.77241 | 5.5915 | 0.77241 | 5.5915 | 16.878 | 2.6 | 3.0639 | 22.1797 | 3.0639 | 22.1797 | 16.15 | 0.070711 | 2.3359 | 0.36846 | 16.9097 | 2.3359 | 16.9097
38 | 12.4302 | 0.20885 | 12.05 | 0.21213 | -0.38018 | -3.0585 | 0.38018 | 3.0585 | 11.7575 | 0.032855 | -0.67267 | -5.4116 | 0.67267 | 5.4116 | 12.3404 | 0.91479 | -0.089787 | -0.72233 | 0.089787 | 0.72233 | 14.8912 | 0.73793 | 2.461 | 19.7989 | 2.461 | 19.7989 | 14.4 | 0.14142 | 1.9698 | 0.29443 | 15.8471 | 1.9698 | 15.8471
39 | 0.72406 | 0.032866 | 0.7295 | 0.006364 | 0.0054357 | 0.75072 | 0.0054357 | 0.75072 | 0.76349 | 0.088111 | 0.039424 | 5.4448 | 0.039424 | 5.4448 | 0.73842 | 0.086048 | 0.014357 | 1.9829 | 0.014357 | 1.9829 | [6 empty] | 0.9 | 0.093338 | 0.17594 | -0.75465 | 24.2984 | 0.17594 | 24.2984
40 | 2.7883 | 0.036048 | 2.82 | 0.014142 | 0.031719 | 1.1376 | 0.031719 | 1.1376 | 3.0371 | 0.41517 | 0.24878 | 8.9222 | 0.24878 | 8.9222 | 2.871 | 0.088085 | 0.082718 | 2.9666 | 0.082718 | 2.9666 | 2.4278 | 0.097197 | -0.36045 | -12.9272 | 0.36045 | 12.9272 | 3.51 | 0 | 0.72172 | -0.14163 | 25.884 | 0.72172 | 25.884
41 | 6.013 | 0.050513 | 5.445 | 0.091924 | -0.56795 | -9.4455 | 0.56795 | 9.4455 | 6.1206 | 1.7861 | 0.10762 | 1.7899 | 0.10762 | 1.7899 | 5.5365 | 0.13068 | -0.47641 | -7.9231 | 0.47641 | 7.9231 | 6.8179 | 0.59533 | 0.80492 | 13.3864 | 0.80492 | 13.3864 | 6.44 | 0.056569 | 0.42705 | -0.36952 | 7.1021 | 0.42705 | 7.1021
42 | 1.8686 | 0.028489 | 1.715 | 0.077782 | -0.15355 | -8.2176 | 0.15355 | 8.2176 | 1.8913 | 0.19265 | 0.022746 | 1.2173 | 0.022746 | 1.2173 | 1.8306 | 0.01799 | -0.037918 | -2.0293 | 0.037918 | 2.0293 | 1.4485 | 0.24299 | -0.4201 | -22.4825 | 0.4201 | 22.4825 | 2.215 | 0.021213 | 0.34645 | -0.46036 | 18.5411 | 0.34645 | 18.5411
43 | 2.503 | 0.0064427 | 2.615 | 0.16263 | 0.112 | 4.4745 | 0.112 | 4.4745 | 3.0529 | 0.076164 | 0.54989 | 21.9693 | 0.54989 | 21.9693 | 2.6036 | 0.017142 | 0.10061 | 4.0197 | 0.10061 | 4.0197 | 2.0842 | 0.072898 | -0.41881 | -16.7323 | 0.41881 | 16.7323 | 3.23 | 0.098995 | 0.727 | -0.13847 | 29.045 | 0.727 | 29.045
44 | 9.4101 | 0.13121 | 9.275 | 0.10607 | -0.13506 | -1.4352 | 0.13506 | 1.4352 | 9.3424 | 0.55704 | -0.067623 | -0.71863 | 0.067623 | 0.71863 | 9.6651 | 0.02444 | 0.25507 | 2.7106 | 0.25507 | 2.7106 | 9.799 | 0.77757 | 0.38891 | 4.133 | 0.38891 | 4.133 | 11.25 | 0.070711 | 1.8399 | 0.2648 | 19.553
45 | 3.6441 | 0.010203 | 3.435 | 0.0070711 | -0.20915 | -5.7393 | 0.20915 | 5.7393 | 3.7752 | 0.10006 | 0.13105 | 3.5962 | 0.13105 | 3.5962 | 3.8815 | 0.071113 | 0.23733 | 6.5127 | 0.23733 | 6.5127 | 4.146 | 0.70468 | 0.5019 | 13.7727 | 0.5019 | 13.7727 | 4.3 | 0.014142 | 0.65585 | -0.1832 | 17.9973 | 0.65585 | 17.9973
46 | 5.1243 | 0.029422 | 4.945 | 0.077782 | -0.17929 | -3.4988 | 0.17929 | 3.4988 | 5.1374 | 0.079151 | 0.013154 | 0.25669 | 0.013154 | 0.25669 | 4.7174 | 0.20842 | -0.40693 | -7.9413 | 0.40693 | 7.9413 | 6.3296 | 0.80386 | 1.2053 | 23.5213 | 1.2053 | 23.5213 | 5.925 | 0.091924 | 0.80071 | -0.096523 | 15.6259 | 0.80071 | 15.6259
47 | 15.1694 | 0.19147 | 14.8 | 0.56569 | -0.3694 | -2.4352 | 0.3694 | 2.4352 | 15.01 | 0.26284 | -0.15942 | -1.0509 | 0.15942 | 1.0509 | 16.069 | 0.5728 | 0.89958 | 5.9302 | 0.89958 | 5.9302 | 17.8402 | 0.1215 | 2.6708 | 17.6065 | 2.6708 | 17.6065 | 17.85 | 0.21213 | 2.6806 | 0.42823 | 17.6711 | 2.6806 | 17.6711
48 | 1.8379 | 0.024443 | 1.915 | 0.049497 | 0.077107 | 4.1954 | 0.077107 | 4.1954 | 1.8786 | 0.12993 | 0.040731 | 2.2162 | 0.040731 | 2.2162 | 1.9119 | 0.02393 | 0.073986 | 4.0256 | 0.073986 | 4.0256 | 1.5172 | 0.34019 | -0.32071 | -17.4499 | 0.32071 | 17.4499 | 2.51 | 0.098995 | 0.67211 | -0.17256 | 36.5694 | 0.67211 | 36.5694
49 | 9.1949 | 0.17769 | 9.01 | 0.18385 | -0.18493 | -2.0112 | 0.18493 | 2.0112 | 9.5357 | 0.37933 | 0.34075 | 3.7058 | 0.34075 | 3.7058 | 9.1714 | 0.35336 | -0.023527 | -0.25587 | 0.023527 | 0.25587 | 10.3402 | 2.2477 | 1.1453 | 12.4555 | 1.1453 | 12.4555 | 11.05 | 0.070711 | 1.8551 | 0.26836 | 20.1749 | 1.8551 | 20.1749
50 | 6.947 | 0.05018 | 6.295 | 0.30406 | -0.65196 | -9.3848 | 0.65196 | 9.3848 | 6.9168 | 0.17921 | -0.030159 | -0.43413 | 0.030159 | 0.43413 | 6.8455 | 0.083672 | -0.10147 | -1.4606 | 0.10147 | 1.4606 | 8.1838 | 0.70468 | 1.2369 | 17.8048 | 1.2369 | 17.8048 | 7.835 | 0.091924 | 0.88804 | -0.051567 | 12.7832 | 0.88804 | 12.7832
SRM 2969 | 1.344 | 0.024 | 1.417 |  | 0.073 | 5.4315 | 0.073 | 5.4315 | 1.674 |  | 0.33 | 24.5536 | 0.33 | 24.5536 | 1.439 |  | 0.095 | 7.0685 | 0.095 | 7.0685 | 1.16 | 0.25 | -0.184 | -13.6905 | 0.184 | 13.6905 | [5 empty] | 0 | 0
SRM 2970 | 1.704 | 0.024 | 1.703 |  | -0.001 | -0.058685 | 0.001 | 0.058685 | 1.853 |  | 0.149 | 8.7441 | 0.149 | 8.7441 | 1.784 |  | 0.08 | 4.6948 | 0.08 | 4.6948 | 1.89 | 0.61 | 0.186 | 10.9155 | 0.186 | 10.9155 | [5 empty] | 0 | 0
SRM 2973 | 7.51 | 0.12 | 7.143 |  | -0.367 | -4.8868 | 0.367 | 4.8868 | 7.331 |  | -0.179 | -2.3835 | 0.179 | 2.3835 | 7.558 |  | 0.048 | 0.63915 | 0.048 | 0.63915 | 8.1 | 0.64 | 0.59 | 7.8562 | 0.59 | 7.8562 | 8.38 | 0.1 |  |  |  | 0 | 0
SRM 972a L2 | 3.39 | 0.26 | 3.22 |  | -0.17 | -5.0147 | 0.17 | 5.0147 | 3.445 |  | 0.055 | 1.6224 | 0.055 | 1.6224 | 3.39 |  | 0 | 0 | 0 | 0 | 3.79 | 0.18 | 0.4 | 11.7994 | 0.4 | 11.7994 | 4.02 | 0.07 |  |  |  | 0 | 0
MEAN |  |  |  |  | -0.090423 | -2.3619 | 0.18658 | 4.4923 |  |  | 0.031137 | 3.7478 | 0.26938 | 7.6046 |  |  | 0.022959 | 1.9249 | 0.24354 | 6.3082 |  |  | 0.79109 | 8.991 | 0.94878 | 14.6163 |  |  | 0.92953 |  | 21.427 | 0.89237 | 21.368
SD |  |  |  |  | 0.2674 | 5.5918 | 0.21054 | 4.0457 |  |  | 0.44319 | 10.4999 | 0.35123 | 8.0975 |  |  | 0.37849 | 8.858 | 0.28847 | 6.4504 |  |  | 0.98197 | 14.3402 | 0.82636 | 8.3331 |  |  | 0.67055 |  | 9.7624 | 0.677 | 9.9897
[5 empty] | CDC | [5 empty] | Univ of WA | [5 empty] | Liege
| NIST |  |  |  | Bias to NIST RMP | [5 empty] | Bias to NIST RMP | [5 empty] | Bias to NIST RMP
| COA Values | Unc. |  |  | Bias | % Bias | ABS BIAS |  |  |  | Bias | % Bias | ABS BIAS |  |  |  | Bias | % Bias | ABS BIAS
SRM2969A-1 |  |  | 1.395 | 0.021213 |  |  |  |  | 1.6104 | 0.079151 |  |  |  |  | 1.3234 | 0.037508
SRM2969B-1 |  |  | 1.415 | 0.049497 |  |  |  |  | 1.6716 | 0.07019 |  |  |  |  | 1.4796 | 0.046673
SRM2969C-1 |  |  | 1.44 | 0.15556 |  |  |  |  | 1.7403 | 0.032855 |  |  |  |  | 1.5145 | 0.0057705
Mean | 1.344 | 0.013 | 1.4167 | 0.022546 | 0.072667 | 5.4067 | 5.4067 |  | 1.6741 | 0.013938 | 0.33011 | 24.5619 | 24.5619 |  | 1.4392 | 0.1018 | 0.09516 | 7.0804 | 7.0804
SRM2970A-1 |  |  | 1.73 | 0.070711 |  |  |  |  | 1.8554 | 0.18668 |  |  |  |  | 1.6936 | 0.10489
SRM2970B-1 |  |  | 1.675 | 0.077782 |  |  |  |  | 1.7868 | 0.227 |  |  |  |  | 1.8436 | 0.18737
SRM2970C-1 |  |  | 1.705 | 0.049497 |  |  |  |  | 1.9177 | 0.17324 |  |  |  |  | 1.8136 | 0.069925
Mean | 1.704 | 0.24 | 1.7033 | 0.027538 | -0.00066667 | -0.039124 | 0.039124 |  | 1.8533 | 0.19564 | 0.14928 | 8.7606 | 8.7606 |  | 1.7836 | 0.079379 | 0.079589 | 4.6707 | 4.6707
SRM2973A-1 |  |  | 7.27 | 0.056569 |  |  |  |  | 7.5578 | 0.022401 |  |  |  |  | 7.582 | 0.33469
SRM2973B-1 |  |  | 6.885 | 0.17678 |  |  |  |  | 7.2389 | 0.094085 |  |  |  |  | 7.4645 | 0.31789
SRM2973C-1 |  |  | 7.33 | 0.028284 |  |  |  |  | 7.1914 | [5 empty] | 7.5071 | 0.10013
SRM2973D-1 |  |  | 7.215 | 0.41719 |  |  |  |  | 7.5631 | 0.28673 |  |  |  |  | 7.6791 | 0.25237
Mean | 7.51 | 0.26 | 7.1433 | 0.23099 | -0.36667 | -4.8824 | 4.8824 |  | 7.3311 | 0.20808 | -0.1789 | -2.3821 | 2.3821 |  | 7.5582 | 0.094106 | 0.04817 | 0.64141 | 0.64141
SRM972aL2A-1 |  |  | 3.22 | 0.042426 |  |  |  |  | 3.4394 | 0.32407 |  |  |  |  | 3.3338 | 0.26188
SRM972aL2B-1 |  |  | 3.215 | 0.0070711 |  |  |  |  | 3.4837 | 0.073177 |  |  |  |  | 3.4 | 0.017142
SRM972aL2C-1 |  |  | 3.23 | 0.042426 |  |  |  |  | 3.1775 | 0.082138 |  |  |  |  | 3.5134 | 0.047861
SRM972aL2D-1 |  |  | 3.215 | 0.14849 |  |  |  |  | 3.6749 | 0.098565 |  |  |  |  | 3.3114 | 0.1536
Mean | 3.39 | 0.12 | 3.22 | 0.0086603 | -0.17 | -5.0147 | 5.0147 |  | 3.4454 | 0.018917 | 0.055376 | 1.6335 | 1.6335 |  | 3.3897 | 0.090675 | -0.00034755 | -0.010252 | 0.010252
|  |  |  | MEAN | -0.11617 | -1.1324 | 3.8357 |  |  |  | 0.088968 | 8.1435 | 9.3345 |  |  |  | 0.055643 | 3.0956 | 3.1007
x =
x =